Form Approved

OMB No. 0920-0307

Exp. Date xx/xx/XXXX

Instructions for Quality Control Assessment - Control Strain Susceptibility Testing

  - Quality Control assessment

    - Quality Control strains:

      - N. gonorrhoeae, strain ATCC 49226 (F-18; Susceptible)

      - N. gonorrhoeae, strain CDC 10328 (intermediate to Ciprofloxacin)

      - N. gonorrhoeae, strain SPL-4 (CMRNG*, CipHLR*, CfxDS*)

*Abbreviations: CMRNG, strain with chromosomally mediated resistance to penicillin and tetracycline (MICs, ≥2.0 µg/ml); HLR, high-level resistance, MIC ≥16.0 µg/ml of ciprofloxacin; CfxDS, decreased susceptibility to cefixime.

    - The LRRB will provide one copy (frozen) of each QC strain annually.

    - ARLN labs are required to propagate and make multiple aliquots of each QC strain upon receiving the isolates.

      - MICs of QC strains should be confirmed at the time the frozen stocks are prepared.

      - Make 5 copies/aliquots for each QC strain.

      - QC strains are stored at -70 oC in a solution of trypticase soy broth containing 20% glycerol.

      - QC strains may be stored at -70 ºC for up to 2 years.

    - QC strains propagation

      - Do not use QC strain that has been passed more than 10 times.

      - Sub-culture QC strains from frozen stock every 7-10 days.

*Notes 1: keep the frozen, stock culture(s) on dry-ice (use ice if dry-ice is not available.)

*Notes 2: work quickly to prevent the frozen, stock culture(s) from thawing.

*Notes 3: return the frozen, stock culture(s) to the -70 ºC immediately.

*Notes 4: discard the stock culture if it had been pulled from the -70 ºC more than 10 times.

    - Quality Control schedule

      - A QC test is performed each time a new batch of growth medium is made/purchased.

      - A QC test is performed each time new batch of antimicrobial agent is used.

      - A QC test is performed each time clinical isolates are tested, unless an Individualized Quality Control Plan (IQCP) is developed that states otherwise.

      - QC results are recorded on QC form or along with test results.

CDC estimates the average public reporting burden for this collection of information as 5 minutes per response, including the time for reviewing instructions, searching existing data/information sources, gathering and maintaining the data/information needed, and completing and reviewing the collection of information. An agency may not conduct or sponsor, and a person is not required to respond to a collection of information unless it displays a currently valid OMB control number. Send comments regarding this burden estimate or any other aspect of this collection of information, including suggestions for reducing this burden to CDC/ATSDR Information Collection Review Office, 1600 Clifton Road NE, MS D-74, Atlanta, Georgia 30333; ATTN: PRA (0920-XXXX).