CDC Model Performance Evaluation Program (MPEP) for

Mycobacterium tuberculosis Drug Susceptibility Testing

Attachment 5

Participant Biosafety Compliance Agreement

Form Approved

OMB No. 0920-6000

Expiration Date: XX/XX/20XX

CDC estimates the average public reporting burden for this collection of information as 5 minutes per response, including the time for reviewing instructions, searching existing data/information sources, gathering and maintaining the data/information needed, and completing and reviewing the collection of information. An agency may not conduct or sponsor, and a person is not required to respond to a collection of information unless it displays a currently valid OMB control number. Send comments regarding this burden estimate or any other aspect of this collection of information, including suggestions for reducing this burden to CDC/ATSDR Information Collection Review Office, 1600 Clifton Road NE, MS H21-8, Atlanta, Georgia 30333; ATTN: PRA (0920-0600).

Model Performance Evaluation Program (MPEP) for Mycobacterium tuberculosis (MTB)

Drug Susceptibility Testing—Participant Biosafety Compliance Letter of Agreement between the  Centers for Disease Control and Prevention and

«Participant Site»         (MPEP «MPEP_ID»)              Year: «XXXX»

Purpose	In the setting of quality management system practices, this letter of agreement is to ensure that the requestor understands and complies with the biosafety guidelines for working with viable strains of Mycobacterium tuberculosis  Complex as outlined in the Guidelines for Safe Work Practices in Human and Animal Clinical Diagnostic Laboratories (MMWR: January 6, 2012).

An authorized representative of Laboratory must sign and date this letter agreement and send by fax or email it to CDC’s Division of TB Elimination Laboratory Branch (DTBE/LB) before Mycobacterium tuberculosis isolates will be sent.  DTBE/LB will only ship isolates to sites meeting applicable biosafety containment. The requestor will be notified if isolates cannot be shipped. Please call 404-639-2862 if you have any questions.

Deadline for submission of this form is (XX/XX/XXXX).

Recommended Laboratory Facilities, Equipment and Practices

Manipulations of viable strains of M. tuberculosis must be performed in a Biosafety Level 3 (BSL-3) laboratory* that meets the following criteria:

- Facility

  - Restricted access to the laboratory with a series of two self-closing doors

  - Seams, walls, and ceiling are sealed and easy to clean and decontaminate

  - Laboratory exhaust air is not re-circulated to any other area of the building

  - A sink is available for hand washing

  - An eyewash station is available

  - The laboratory is equipped with a Class II-A1 or II-A2 Biological Safety Cabinet (BSC) that is re-certified annually

- Standard Microbiological Practices

  - A risk assessment for working with M. tuberculosis in the laboratory has been completed

  - Biosafety is addressed in laboratory protocols using the requested organism

  - Laboratory personnel are knowledgeable about hazards of working with M. tuberculosis and demonstrate proficiency in laboratory procedures

  - BSL-3 personal protective equipment (PPE) is used based on risk assessment of laboratory procedures and practices

  - All procedures involving manipulation of viable strains of M. tuberculosis are conducted in a BSC or other physical containment devices

  - Decontamination of laboratory waste must be available in the facility, preferably within the laboratory (e.g., autoclave, chemical disinfection, or other validated decontamination method)

*If a BSL-3 laboratory is not available, a BSL-2 laboratory may be used if the following requirements are met:

- A risk assessment determines that viable strains of M. tuberculosis can safely be worked with in the separate, closed BSL-2 laboratory using BSL-3 practices as outlined in the above mentioned MMWR supplement

- Laboratory exhaust air is exhausted to the outside of the building

- The laboratory director approves the practice

_________________________________________________________________________________________

Participant Compliance						(MPEP «MPEP_ID»)

Please initial the statement that applies to your laboratory:

___	Meets all of the above recommendations for a BSL-3 laboratory

___  	No BSL-3 laboratory is available but meets all of the laboratory requirements as described above for a BSL-2 laboratory with BSL-3 practices

___ 	No BSL-3 laboratory is available and does not meet all of the laboratory requirements as described above for a BSL-2 laboratory

Laboratory warrants and represents that the information provided above is current, complete and accurate. The Laboratory acknowledges and understands the potential risks associated with manipulation of live cultures of M. tuberculosis and will work with these isolates under the appropriate biosafety containment as determined by our institutional policies and procedures. Use of the isolates and derived material will be confined to quality management systems practices, to include verification, validation, panel and proficiency testing, and not research and development

___________________________________	    _______________________________________

Person Completing Form (Print)			       Position/Title

___________________________________	    _______________________________________

Person Completing Form (Signature)	  		Name of Organization

________________				_______________________________________

Date								City/State

Please email a PDF attachment of the signed document or fax the signed document to:

Laboratory Branch (TB) MPEP Program

CDC MPEP email: TBMPEP@cdc.gov

Fax: 404-639-5491